STERILE PROCESSING AUDIT TOOL

| Date: | Audit conducted by: |
| --- | --- |
| Personnel Considerations | - Orientation and education performed and documented - Demonstrated knowledge of and documented competence - Continuing education at regular intervals - Training for all new instrumentation, devices, and equipment - Consistently adhering to dress code - Facility laundered scrub attire donned at hospital - All facial hair covered (except for eyebrows) - PPE used and removed properly - Manufacturers’ written instructions for use (IFU) available and followed - Supervisors - Certified in Sterile Processing Management - Demonstrate comprehensive understanding of - Relevant state and federal regulations - Occupational Safety and Health Administration (OSHA) blood borne pathogens exposure control plan - Engineering and work-practice controls - Actively participate in committees such as: - Infection Prevention and Control - Quality improvement (sterilization reports sent to IP) - Safety, and - Product Evaluation and standardization - Staff certified within 2 years |
| Facility Design Considerations | - Space proportioned to expected volume - Floors and walls level and can withstand frequent cleaning - Ceilings flush surface, no shedding materials with recessed and enclosed pipes - Doors and windows kept closed - Handwashing stations conveniently located in clean and decontamination areas - Eye wash stations located within 10 seconds travel time – 15 minute flush - Equipment maintenance records maintained - Functional work flow pattern - Dirty to clean - Pass-through windows and all doors kept closed - Traffic pattern - Restricted to authorized personnel only - Signs posted - Cardboard boxes removed before items are brought to clean areas - Temperature and humidity levels monitored and recorded daily - Temperature - 68-73° F clean areas - 60-65° F decontamination areas - Humidity - 30-60% in work areas - Not over 70% in sterile storage areas - Ventilation - Soiled area, negative – 10 air exchanges per hour - Clean/sterile area, positive – 10 air exchanges per hour - Water quality matches manufacturers’ recommendations - Distilled, deionized or reverse osmosis water use for final rinse on instruments - Housekeeping procedures - All areas cleaned daily (should be the same as in OR) - Separate cleaning equipment for decontamination |
| Decontamination | - Appropriate decontam PPE available - Hands are washed after removing PPE - Manufacturers’ written instructions for use (IFU) available and followed - Equipment - Instruments - Cleaning solutions - Appropriate cleaning and decontamination solutions - Dilution – measuring cups and lines in the sink for accurate measuring - Expiration dates - Solution containers labeled - Appropriate cleaning processes used - Sharps and delicates separate - No use of saline on instruments - Cleaning happens as soon as possible (point of use) - Instruments kept moist until they are cleaned - Not cleaned in hand sinks or scrub sinks - Brushing occurs under water - Brushes are disposable or decontaminated daily - Water quality meets manufacturers requirements - Disinfectant concentration is tested per manufacturers’ recommendations - Contaminated items transported in covered containers marked biohazard |
| Packaging | - Labeling on indicator tape, patient record cards or plastic side of peel packs - Instrument set weights not over 25 pounds - Peel packs – double peel packs are not folded, proper size used - Internal and external chemical indicators (CI) used for all packages - Geometric center of wrapped packages - Rigid containers – Follow manufacturers’ written recommendations - On all levels - Instruments are in good condition - Instrument refurbishing plan - Instrument tape (if used) is in good condition - Tip protectors are validated for use - Any Single Use Devices reprocessed (needs to be FDA cleared) - Instrument tracking system available |
| Sterilization | - Loading and unloading practices - Peel packs and lighter items on top shelf - Peel packs and linen packs are set on edge (not horizontal) - No stacking of pans (without manufacturers’ recommendations) - Documentation for each load - Sterilizer identification - Type of sterilizer and cycle used - Lot control number - Load contents - Critical parameters for specific sterilization method - Operator’s name, and - Results of the sterilization process monitors (physical, CI, BI) - Sterilization Monitors - BI – run daily in steam and peracetic acid - BI – run in every load for ethylene oxide, gas plasma, vapor phase, or ozone - The same lot number is used for the control and the processed BI - Sterilization records storage follows the facilities’ record retention policy - Current manufacturers’ recommendations readily available and followed - Extended cycles run per manufacturers’ recommendations - Implants - Monitored with a biological indicator (BI) and a class 5 CI - Not released until results of BI available - Management of loaner instrumentation program in place - All sterilized items are traceable to the patient - Recall process in place and reported to Infection Prevention and Control - Flash sterilization practices - Items are appropriately cleaned - Use of closed validated flash containers - All parameters documented - Aseptic transportation to point of use - Implants - BI and CI run with all implants - Not released until results of BI are available - Traceable to the patient - Not used as a substitute for sufficient instrument inventory |
| Sterile Storage | - Storage conditions - Cleanable surfaces - Bottom shelves are solid - All items are - 18” below the ceiling (or level of sprinkler head) - 8-10” above the floor - 2” for outside walls - Sterile items separate from clean items - Heavy wrapped trays are not stacked - Shelf life/event related – stock rotation - Controlled area (appropriately attired persons only) signs posted - Handwashing facilities readily available - No web-edged or corrugated boxes |
| Policies and Procedures | - Policies are updated according to current best practices - All sterilization and reprocessing policies and procedures readily available for staff - Dress code followed (no artificial nails or polish) - Care and handling of instruments and powered equipment - Packaging systems – selection and use - Sterilization recall - Sterile storage - Chemical disinfectant (including high-level disinfecting) - Shelf Life (event related) - Preventive Maintenance for equipment - Steam shutdown - Sterilization – steam and low temp - Endoscopes – cleaning and processing - Environmental cleaning - Creutzfeldt-Jakob disease (CJD) - Toxic Anterior Segment Syndrome (TASS) - Management of Loaner Instrumentation - Single Use Devices - Up-to-date and referenced - Followed and monitored - Audits completed |
| Comments |  |